Validation of a LC method for the determination of levonorgestrel released from a polymeric matrix.

A two days protocol.

C. Hubert1, E. Rozet1, B. Boulanger2, W. Dewé2, A. Breuer3, L. Collard3, B. Evrard3, Ph. Hubert1

1Laboratory of Analytical Chemistry, Bioanalytical Chemistry Research Unit, Institute of Pharmacy, University of Liège, B 36, 4000 Liège, Belgium, ²Lilly Development Centre, 1348 Mont-Saint-Guibert, Belgium, 3Laboratory of pharmaceutical technology, Institute of Pharmacy, University of Liège, B 36, 4000 Liège, Belgium

The present LC method was developed for the determination of levonorgestrel amount released from a polymeric matrix in a dosing range covering 30 ng/mL to 1000 ng/mL. The dissolution medium selected for that purpose was aqua conservans (0.1 % parabens). Nowadays, the analytical methods must not only be fully validated but also in a time as short as possible. In this paradoxical situation a new regulatory request is to manage the risk associated to the use of these methods in routine analysis (1-2). In this context, the accuracy profile concept using an appropriate experimental design turns out to be a very useful tool (3). In the present study, a validation protocol including simultaneously equipments, operators, days and preparations of stocks solutions as variability sources was elaborated. Being without knowledge about a potential matrix effect, two calibration curves (within and without matrix) were also investigated. On the basis of the different accuracy profiles obtained, a very significant matrix effect was highlighted. Finally, the calibration was performed within the matrix using a weighted linear regression model. Excellent reliability was obtained. Indeed, relative bias and intermediate precision were lower than 1.5 % and 4 %, respectively. Moreover, the risk to obtain a result outside of the acceptance limits (15 %) is lower than 3 %. All these results were conducted in only two days.

(1)	PAT Initiative, FDA, http://www.fda.gov/cder/OPS/PAT.htm

(2)	ICH guidelines Q9, EMEA, http://www.emea.eu.int/Inspections/docs/ICHQ9Step4QRM.pdf

(3)	Ph. Hubert et al, Harmonization of strategies for the validation of quantitative analytical procedures: A SFSTP proposal—part I, J. Pharm. Biomed. Anal. Vol. 36, 579 (2004)